Signet 4150 Engineering Specification

- The controller shall be a microprocessor-based instrument with two-line backlit LCD.

- The turbidimeter shall meet all requirements specified by the USEPA Method 180.1.

- Shall be available in ranges from 0-100 NTU and 0-1,000 NTU.

- Accuracy shall be 2% of reading or +/- plus or minus .02, whichever is greater, from 0-40 NTU, and 5% of reading or +/- plus or minus .02, whichever is greater, from 40-1000 NTU. Resolution will be 0.0001 NTU (selectable).

- The turbidimeter enclosure shall be NEMA 4X (IP66).

- The turbidimeter shall be Modbus compatible with 4-20mA selectable

- The turbidimeter shall have two 120-240 VAC, 2A SPDT Form C Relays

- The turbidimeter shall be certified ETL and CSA 22.2 No. 1010-1-92

- The continuous monitoring system shall include a single modular unit with 100 – 240 VAC, 50 – 60 Hz power supply, display and sensor as one single unit.

- The turbidimeter shall have the option of using an automatic ultrasonic cleaning system in finished or raw water applications.

- The sensor shall consist of a rotational flow through assembly with a 30ml cuvette. The specially designed flow head bubble rejection system eliminates the need for a bubble trap and ensures an immediate response time.

- The sensor shall be able to accommodate grab samples.

- Calibration and standardization will be accomplished using small volumes (30ml) of reusable primary standards in a cuvette. The Primary Standards shall be reusable for multiple online turbidimeters and be interchangeable with laboratory turbidimeters

- Calibration procedures can be completed without disrupting the sample flow.

- The lamp source and detector shall not come in contact with the sample, eliminating false low readings.

- The Online Turbidimeter shall be a Signet Model 4150.