Specification For pH/ORP Electrodes

- The pH or ORP electrode mounting options shall include inline via ¾" MNPT process connection or submersible.

- The pH electrode shall have built-in temperature compensation, utilizing 3k ohm balco, or PT 1000, and ORP shall utilize a 10 k ohm ID resistor for self-configuring transmitters

- The pH electrode shall have a measurement range of 0–14pH, and ORP electrode shall have a measurement range of +/- 2000 millivolts.

- The pH or ORP electrodes shall be capable of being mounted at any position around a horizontal or vertical pipe.

- The pH or ORP electrodes shall have an operating temperature range of 0°C–85°C (32°F–185°F), with a maximum operating pressure range of 100 psi.

- The pH electrode wetted materials shall be Ryton (PPS - PolyPhenylene Sulfide) body, glass sensing surface, and PTFE reference junction, with FPM seals.

- The ORP electrode wetted materials shall be Ryton (PPS) body, platinum  sensing surface, and PTFE reference junction, with FPM seals.

- The pH or ORP electrodes shall utilize DryLoc connection pH/ORP pre-amps and blind transmitters.

- The pH electrodes shall be Georg Fischer Signet model 2774 flat glass or equal, or 2776 bulb or equal; ORP electrode shall be a 2775 bulb or 2777 flat or equal.